## Slide 1: TPAC1.1 Testing JC/Jan 16th

## Slide 2: Comparator Investigations
Two (related) symptoms were observed
  Non-gaussian threshold scans, with steep sides and flat tops
  Test pixel comparator oscillates for very low thresholds
Coupling (positive feedback) in the pixel design
  Not present (or significantly diminished) in old pixel design
  Dominant in new pixel design causing oscillations
  Schematic is fundamentally the same, so must be due to layout  parasitic capacitance
Full RCX simulation
  Fixed a bug (previously caused full pixel RCX to fail)
  Setting a very low threshold shows oscillations!
    Good  any fix should be possible to prove in simulation
    Previously all RCX simulations had not checked the very low threshold case, which is necessary to cause oscillation
    Pixel performs ok at higher thresholds, although some injection can be seen that might have identified a potential problem
  Doesn’t actually identify the problem, just confirms that one exists
Analysis of circuit and observed behaviour is required to understand what is going on

## Slide 3: Comparator Circuit Analysis
Polarity of injection eliminates coupling between certain nodes
Eventually found that a single parasitic capacitance between comparator output and diode node can cause oscillations at low thresholds
  RCX extracts 30aF between these two nets in the 1.1 pixel design (v small!)
  RCX extracts no parasitics between these two nets in the 1.0 design
  Schematic simulation (no parasitics) with an additional 30aF between the two critical nodes shows oscillations at low thresholds.

## Slide 4: Sanity check
Can 30aF really matter?
  Would not normally consider such tiny parasitics!
  But… diode node is sensitive to induced charge, with a large gain…
Consider a switching 1.8v signal coupling through a 30aF capacitor…
Q = 30 x10-18 * 1.8 = 5.4 x10-17 C
= 337 electrons
Circuit charge gain is ~140uV/e- so… = 47mV signal
!

## Slide 5: TPAC1.0 preShape pixel layout
Comparator output
M1 M2 M3 M4
Comparator output bridges diode node only once, on metal 4 with metal 3 shield.

## Slide 6: TPAC1.1 preShape pixel layout
Comparator output
M1 M2 M3
Comparator output was re-routed in v1.1 over SRAMS but crossing diode node twice
No M2 shield at one crossing creates dominant capacitance between the two nets

## Slide 7: TPAC1.2 preShape pixel layout
Comparator output
M1 M2 M3
Single mask change (M2)
Extended shielding (ground)
RCX tool finds no parasitics between comparator output and diode node 
?

## Slide 8: Simulation Summary
Design | View | Cpara (HIT DIODE) | Cpara (DIODE GND) | Gain μV/e- | Simulation
Original 1.0 | Schematic | 0 | 14 (est) | 118 | 
 | Extracted (C only) | 0 | 13 | 164 | 
 | Extracted (sel RC) | 0 | 13.3 | 164 | 
Revision 1.1 | Schematic | 0 | 14 (est) | 136 | 
 | Extracted (C only) | 30.25a | 12.1 | 182 | Oscillates at low Vth
 | Extracted (sel RC) | 27.9a | 12.4 | 181 | Oscillates at low Vth
Amended 1.2 | Schematic | 0 | 14 (est) | 160 | 
 | Extracted (C only) | 0 | 12.5 | 180 | 
 | Extracted (sel RC) | 0 | 12.8 | 178 | 

## Slide 9: Comments
Suggested fix adds small additional parasitics to diode node, but acceptable within context of original design
Unsure of reliability of parasitic extraction tools at this precision (10-18)
  what error bars to apply?
Small injection effects are seen in the v1.0 test pixel
  which the RCX tool does not predict
  but the pixel does operate properly

## Slide 10: Measured Injection
Cross check size of injected signal with predicted coupling capacitance
  Charge gain known from marcel’s 55Fe test pixel results
  Can observe signals at two points in analog chain
Induced signal on shaper output
  Varies, in range 1724mV
    Note: Corresponding injection on shaper input will be too small to see on scope (<1mV)
  Applying a gain of 150uV/e-
  Injection varies in range 113160e-
    Right order of magnitude! 

## Slide 11: Mask change costs
M2 required to fix parasitic capacitance
CS required to fix address repettition
$$…

## Slide 12: Spare slides
Signals during oscillation
  Triggered by noise
  Oscillation
  Similar scope trace

## Slide 13: Signals
Diode
Preamp out
Shaper out
Comparator
Threshold

## Slide 14: Signals
Diode
Preamp out
Shaper out
Comparator
Threshold

## Slide 15: Signals
Diode
Preamp out
Shaper out
Comparator
Threshold

## Slide 16: Signals
Diode
Preamp out
Shaper out
Comparator
Threshold

## Slide 17: Signals
Diode
Preamp out
Shaper out
Comparator
Threshold
etc…

## Slide 18: Signals
Diode
Preamp out
Shaper out
Comparator